**MyScreen Application Note**

**For use with the Illumina® MiSeq™/ MiniSeq™**

**For Research Use Only (RUO)**

|  |
| --- |
| **Manufacturer:**  **Developed by Gamidor Diagnostics Ltd. Manufactured by © Agilent Technologies, Inc. 2017-2019**  **Version C2, July 2019 Printed in USA Agilent Technologies, Inc. 5301 Stevens Creek Blvd Santa Clara, CA 95051 USA**  **Technical Support**  **Call 00800 345 600 (toll free) or +49 69 8679 7730 Or send an e-mail to:** [**genomics\_tech\_europe@agilent.com**](mailto:genomics_tech_europe@agilent.com) |
| **Distributed by:**  **Gamidor Diagnostics Ltd.**  **Hasaham St. 32 Petach Tikva , Israel**  **Phone: 03-9277277**  [**www.gamidor.com**](http://www.gamidor.com) |
| **For Tech Support:**  [**amichai@gamidor.co.il**](mailto:amichai@gamidor.co.il)  **052-2808772**  **For Professional use only, for more information** <https://gamidor.co.il/myscreen/> |

Table of Contents

Introduction 3

I. Intended Use 3

II. Background 3

Protocol User Guides 4

Precautions 5

System and software guides from Illumina® 6

Sequencing runs performance 7

Analyzing the Data 8

A. Generating analysis files 8

B. Viewing the result 9

C. OPTIONAL - VariantStudio3.0 11

D. Viewing Raw Sequencing data using IGV. 11

E. Verifying a BIG DELETION using IGV: 12

Appendix A: Software Installations Guides 13

F. Dependencies 13

G. Supplemental Software 13

H. Installation 14

Appendix B: Appendix Directories 15

I. Genotyping Error! Bookmark not defined.

J. DECoN Error! Bookmark not defined.

K. Patient Reports Creator Error! Bookmark not defined.

Appendix C: Version Update Protocol 15

Appendix D: Legal Notices 15

L. Methods and Limitations 15

Introduction

## Intended Use

The MyScreen Kit is used to detect and identify a panel of hundreds of different genetic diseases that are associated with various genetic disorders of the Israeli population. It is covering most of Israeli ethnic groups including Jewish, Arabs, Bedouins, Christians, Druze.

The panel covers all Jewish ethnic groups including Ashkenazi and Non-Ashkenazi (Morocco, Libya, Algiers, Turkey, Bulgarian, Iraq, Iran, Egypt, Syria, Lebanon, Bukhara, Kurdistan, Yemenite etc.).

The panel list was designed with data collected from:

* + The Department of Community Genetics at the Ministry of health - MOH (סל הבריאות)
  + Israeli Society of Medical Genetics –ISMG (איגוד הגנטיקאיים)
  + The Israeli National Genetic Database (Golden Helix)
  + Pathogenic mutations described in Israeli populations with unknown prevalence that were gathered from several sources such as: medical geneticists, medical teams, publications and more.

**Masking of any diseases from the panel is possible following MOH or ISMG request.**

***For Research Use Only. Preforming this test should be done by a professional lab technician who have been trained in NGS technologies and qualified by Illumina representative.***

## Background

Israeli population comprised of many ethnic backgrounds. Different ethnic groups have a higher risk for specific disease-causing mutations than the general Israeli population. These diseases are inherited in an autosomal or x-linked recessive pattern.

The MyScreen assay is a targeted assay for a selected set of mutations per disease and per specific genes. The list of chosen diseases and selected mutations might be updated from time to time.

The detailed list of all mutations detected in this assay can be requested from: [amichai@gamidor.com](file:///C:\Users\hadas\Downloads\amichai@gamidor.com)

Protocol User Guides

**SureSelect XT HS and XT Low Input Enzymatic Fragmentation Kit. Agilent Technologies**

<https://www.agilent.com/cs/library/usermanuals/public/G9702-90050.pdf>

**SureSelectXT HS Target Enrichment System for Illumina Paired-End Multiplexed Sequencing Library. Agilent Technologies**

Instructions for preparing samples using the SureselectXT HS Target Enrichment System

<https://www.agilent.com/cs/library/usermanuals/public/G9703-90000.pdf>

**MiSeq™ System Denature and Dilute Libraries Guide**

This guide provides instructions for denaturing and diluting libraries after llibrary preparation and before sequencing on the MiSeq™. This guide also includes instructions for preparing a PhiX control.

<https://support.illumina.com/content/dam/illumina-support/documents/documentation/system_documentation/miseq/miseq-denature-dilute-libraries-guide-15039740-10.pdf>

**MiniSeq™ System Denature and Dilute Libraries Guide**

This guide contains instructions for denaturing and diluting libraries before sequencing on the MiniSeq™ System, and instructions for preparing a PhiX control.

<https://support.illumina.com/content/dam/illumina-support/documents/documentation/system_documentation/miniseq/miniseq-denature-dilute-libraries-guide-1000000002697-00.pdf>

**Technical Assistance**

Specialists from the Technical Assistance Center can help troubleshoot and resolve problems. Contact the Center via one of the following methods:

|  |  |
| --- | --- |
| Phone: | 052-2808772 |
| Email: | [amichai@gamidor.co.il](file:///C:\Users\hadas\Downloads\amichai@gamidor.co.il) |

Precautions

• To prevent contamination of reagents by nucleases, always wear powder- free laboratory gloves and use dedicated solutions and pipettors with nuclease- free aerosol- resistant tips.

• Use best-practices to prevent PCR product contamination of samples throughout the workflow:

1. Assign separate pre-PCR and post- PCR work areas and use dedicated equipment, supplies, and reagents in each area. In particular, never use materials designated to post- PCR work areas for pre- PCR segments of the workflow.
2. Maintain clean work areas. Clean pre-PCR surfaces that pose the highest risk of contamination daily using a 10% bleach solution, or equivalent.
3. Always use dedicated pre-PCR pipettors with nuclease- free aerosol- resistant tips to pipette dedicated pre- PCR solutions.
4. Wear powder-free gloves. Use good laboratory hygiene, including changing gloves after contact with any potentially- contaminated surfaces.

• For each protocol step that requires removal of tube cap strips, reseal the tubes with a fresh strip of domed caps. Cap deformation may result from exposure of the cap strips to the heated lid of the thermal cycler and from other procedural steps. Reuse of strip caps can cause sample loss, sample contamination, or imprecision in sample temperatures during thermal cycler incubation steps.

• In general, follow Biosafety Level 1 (BL1) safety rules.

• Possible stopping points, where samples may be stored at –20°C, are marked in the protocol. Do not subject the samples to multiple freeze/thaw cycles.

**References for Contamination Control:**

Kwok, S. and Higuchi, R. (1989). Avoiding false positives with PCR. Nature (London) 339, 237.

Victor, T. et al. (1993). Laboratory experience and guidelines for avoiding false positive polymerase chain reaction results. Eur. J. Clin. Chem. Clin. Biochem. 31, 531.

Yap, E.P.H. et al. (1994). False-positives and contamination in PCR. In: PCR Technology: Current

Innovations. Griffin, H.G. and Griffin, A.M., eds., CRC Press, Boca Raton, FL.

System and software guides from Illumina®

**MiSeq™ System User Guide**

This guide includes information about instrument components, MiSeq™ Control Software (MCS), and required consumables, and instructions for performing a sequencing run and maintaining the instrument.

<https://support.illumina.com/content/dam/illumina-support/documents/documentation/system_documentation/miseq/miseq-system-guide-15027617-01.pdf>

**MiSeq™ Reporter Analysis Workflow Reference Guides**

Each workflow reference guide contains an overview of the analysis workflow and descriptions of analysis output files. For information about the MiSeq™ Reporter interface and instructions for instructions for installing, using, and troubleshooting the software

<https://emea.support.illumina.com/content/dam/illumina-support/documents/documentation/software_documentation/miseqreporter/miseq-reporter-enrichment-workflow-guide-15042315-01.pdf>

**MiniSeq™ System Guide**

This guide contains an overview of instrument components and instructions for operating and maintaining the

MiniSeq™ System

<https://support.illumina.com/content/dam/illumina-support/documents/documentation/system_documentation/miseq/miseq-system-guide-15027617-01.pdf>

Sequencing runs performance

1. When the Miseq™/Miniseq™ run ends copy the following files from output folder \***NAME OF RUN***\ to SAV (Sequencing Analysis Viewer) analysis:

* runinfo.xml file,
* runparameters.xml file
* InterOp folder

1. Download the Sequencing Analysis Viewer (SAV) software from this website: <https://support.illumina.com/sequencing/sequencing_software/sequencing_analysis_viewer_sav/downloads.html>
2. Install the SAV software on your computer and go over your run parameters by following the instructions in the user guide you can find in this link:

<https://support.illumina.com/sequencing/sequencing_software/sequencing_analysis_viewer_sav/documentation.html>

1. For MiSeq™, you can find the recommended specifications regarding reads or cluster passing filter, cluster generation, sequencing and quality of the data in this link:

<https://www.illumina.com/systems/sequencing-platforms/miseq/specifications.html>

For MiniSeq™, you can find the recommended specifications regarding reads or cluster passing filter, cluster generation, sequencing and quality of the data in this link:

<https://www.illumina.com/systems/sequencing-platforms/miniseq/specifications.html>

Analyzing the Data

## Generating analysis files

1. Once the run and Illumina® analysis completed, copy (using Disk on Key) the following files for MyScreen analysis under the output run folder in

* LRM (Local Run Manger) system in MiniSeq \***NAME OF RUN***\Alignment\_1\xxx\_xxx\:
* \*.bam
* \*.bai
* aggregate.report.pdf - exists only in LRM system
* MiSeq Reporter - Miseq ***NAME OF RUN*** \Data\Intensities\BaseCalls
* \*.bam
* \*.bai

***NAME OF RUN*** - for example 191216\_M06216\_0016\_000000000-G4NJR

1. Save all those files in a new results library – COPY THE ***NAME OF RUN*** in your analysis computer (dedicated to MyScreen results).
2. Quality Control Step –

Open “aggregate.report.pdf”, on page 7 look at the “Coverage Summary” table and verify:

* Mean Region Coverage Depth above 150 per sample
* Uniformity of Coverage above 90% per sample

If any of the above does not meet these thresholds you need to consider creating a new library for this sample. Rare analytic errors may occur that interfere with reporting.

1. In order to activate the MyScreen program analysis, double-click on the file named "MyScreen.exe”. A prompt interface will guide you to **Run Analysis** or to review version’s mutations list and supported files. These files are version dependet and can be located under “MyScreen\נספחים ופרוטוקולים”:

* MyScreen.V2.#\_App.Note – Appliction Note for current version
* MyScreen\_VALIDATED-Postive\_(*date*)Ver2.#.xlsx – Excel file aggregating positive samples detected by MyScreen.
* (*date*) Ver 2.# - רשימת גנים ומחלות.xlsx – List of genes and diseases discovered in current MyScreen version.
* (*date*) Ver 2.# - רשימת מחלות ומוטציות גירסה – List of mutations detectable in current version.

**Run Analysis**

1. Choose the directory containing bam and bai files.
2. Choose detection panel for all samples or per sample from the drop-down menu.
3. Select Extra Information Excel file to add to produced reports (optional). Sheet1 of the Excel file must be defined the personal information of each sample (like ID, name, address and more) to create final report (PDF and WORD) for each sample.
4. Insert the ***NAME OF RUN***. By default, the directory name in step 1 will be written. This name must contain hospital’s MiSeq/MiniSeq machine code (####\_M…).
5. Select if Word and PDF reports are needed for current run. If yes, choose hospital template for reports.

**Create Summaries**

In case you wish only to produce results files and reports for an already analyzed run.

The software will start process the genotyping + CNV analysis (should takes around 10 min for sample) and will prompt a successful message:

“Genotyping Analysis Completed!”

CNV detection completed!

Word reports created

Converting reports to pdf...

Press any key to continue . . .

## Viewing the result

A new library will be created under ***NAME OF RUN*** named MyScreen\_Analysis\_v#\_RESULTS

1. The file “sample\_summary\_[DATE]” contains a summary of all the samples and their results (wild type or respective mutations from the current run).
   1. Each sample will be output as a result.
   2. If no mutations or problems were found it will be indicated as “No mutations identified (WT)”
   3. If genotyping/CNV mutations are found, the mutations will be detailed with its output parameters.
   4. If problem/no-calls mutations are found, the mutations will be highlighted in yellow for problem. Every problem should be checked by the lab using IGV or other alternative method to verify the heterozygotes status.
2. All reports have been created in a directory named MyScreen\_Analysis\_v#\_RESULTS in a folder named REPORTS.
   1. All reports named in the following format: [sample name]\_[ID].docx/.pdf
   2. Only validated positives from our run will be appeared in the REPORTS.
   3. Positives indicated as Problems will be appeared as “Low Confidence” in the REPORTS
3. The file “db\_statistics\_[DATE]” containing ALL sample results obtained from ALL past runs performed (including the current run) - this file can be used for statistics.
4. ‘errors.log’ file will be created in the main analysis output directory, if any variants failed to be detected in the analysis or any failed sample detected in the CNV analysis.

**OPTIONAL - Info directory in the MyScreen\_Analysis\_v#\_RESULTS that can viewed:**

1. **Results.xlsx –**

All information regarding the genotyping analysis, sheet separated to cases; Geno Full, Geno Positive, Geno WT Poly, Geno Gender, Geno NonReported, CNV Calls, CNV Failed Samples, CNV Fail Exon.

1. CNV Calls – Contain list of CNV calls made by DECoN with the appropriate parameters associate to each call.

Low coverage mutations or problematic Alternative Variant frequency variants will be displayed as “CNV-Problem”.

The following columns in the file are:

* Type – Deletion (Duplication in DMD only)
* Custom.first and Custom.last – Exons deleted in gene
* Correlation – Sample correlation in sample set
* N.Comp – Number of samples compared with in the anlysis
* BF – Bayes Factor, statistical strength of evidence in favor of one theory among two competing theories. > 10 strong evidence, less than 10 = CNV-Problem (not reported)
* Info – If a sample did not pass quality control

1. CNV Failed samples – If a sample did not pass minimum correlation threshold it will be listed in current sheet.
2. CNV Failed exons – Listing gene exons that fail to pass minimum read coverage limit.
3. **MyScreen\_Analysis-[*Date*].log** and **.out.log**

Log files summarizing the current run. Must be included when contacting support.

1. **AG\_DB\_[*Date*].pbz2**

Compressed results aggregated file of current run. For internal use to reproduce run results.

1. **Genotyping Directory –**

* Intermediate files and logs used during analysis run.
* \*Optional - VCF to load to the VariantStidio3.0 files

1. **CNV Directory –**

* Intermediate files and logs used during analysis run.

## OPTIONAL - VariantStudio3.0

Basic recommended steps to view analyze and report data with VCF files

To view the data on the VariantStudio use Illumina® [VariantStudio User Guide](https://support.illumina.com/content/dam/illumina-support/documents/documentation/software_documentation/variantstudio/variant-studio-v3-0-user-guide-1504890-01.pdf)

**Result viewing:**

1. Open VariantStudio 3.0
2. Open MyScreenProject.VSProj under “C:\Gamidor\Appendix” (or from the recent menu option)
3. Load vcf files from the Home tab “Import VCF” or “Import Folder” under Info/Genotyping
4. Check “Load hom-ref positions”
5. Click OK
6. ManageSamples - From the Home tab, use commands on the Samples menu to import variant call files in VCF file format and manage samples in the project.
7. From the Home tab, select “View All Samples”. It will show all samples in the project.

**Result exporting:**

To export text/csv file: VariantStudio provides tools for exporting to text files and graphical representations of data. From the Export menu in the Reports tab, select an option to “Filtered variants” data to a tab-separated values file. These text file formats are not application-specific and can be opened in any text editor.

## Viewing Raw Sequencing data using IGV.

1. Double click on IGV.bat to start the program.
2. Genomes are selected from the genome drop-down list on the upper-left of the IGV window. Choose hg19.
3. To load data from the file system: (The preferred file format for viewing alignments in IGV is the BAM format.) Load data files by browsing for files on the local file system.
4. Select File>Load from File. IGV displays the Select Files window.
5. Select one or more data files or sample information files, then click OK.
6. IGV will display a warning if the file is an un-indexed ASCII-format file over 50 MB. It is recommended that such files should be indexed or converted to the binary TDF format prior to loading (see section on igvtools). Meaning BAM and BAI files should be located in the same directory.
7. Use the search box to locate: A locus or a mutation (for example, chr5:90,339,000-90,349,000), A gene symbol or other feature identifier (e.g., DPYD or NM\_10000000)
8. When searching features, IGV will accept partial matches. A mutation in a feature. IGV accepts 2 mutation formats: Amino acid mutation notation of this form:
9. KRAS:G12C - The above would look for a mutation in KRAS on the 12th amino acid, from Glycine to Cystine. \* is stop codon
10. Nucleotide mutation of this form: KRAS:123A>T - The above would look for a mutation in KRAS from adenine to thymine at the 123rd base.


For more details please use IGV user guide:

<http://software.broadinstitute.org/software/igv/book/export/html/6>

## Verifying a BIG DELETION using IGV:

1. Open the positive sample with at least 2 samples with similar mean coverage.
2. Use the search box to locate the big deletion boundaries (from the samples\_summary,xls in the reports, for example, chr5:90,339,000-90,349,000)
3. You need to verify all the following

* There is no polymorphism or mutations in the deleted region, which could imply 2 alleles.
* Search the Break point region by reduction of the coverage region
* Look on parallel regions and verify that the number of reads in on average is half for the positive sample versus the controls ones



Appendix A: Software Installations Guides

A dedicated computer will be allocated for analysis.

The computer does not have to be connected to the Internet.

It is recommended that it will be connected to the MiSeq/MiniSeq tp transfer the results easily.

- computing Requirements for the computer:

* ≥ 16 GB RAM
* ≥ 1 TB of disk space
* 64-bit quad core processor (2.8 GHz or higher) with the AVX instruction set (recommended) or Intel Core
* i7-210QE 2.10 GHz (or equivalent) processor with the AVX instruction set (minimum).  
   Display resolution of 1024 x 768 or higher

Software Requirements

* Windows 7 Pro 64-bit with Service Pack 1 (English-US) or Windows 10 (Personal Edition versions are not supported)
* Office 2016

## Dependencies

* Python version 3.6/3.7/3.8: Added to Environment Variable Path as python

DOWNLOAD [here](https://www.python.org/downloads/)

* JAVA (JRE and JDK): Main directory "Java" lies in "C:\Program Files".

DOWNLOAD JRE [here](http://www.oracle.com/technetwork/java/javase/downloads/jre8-downloads-2133155.html)

DOWNLOAD JDK [here](http://www.oracle.com/technetwork/java/javase/downloads/jdk8-downloads-2133151.html)

* Visual Studio 2017 version 15.9:

DOWNLOAD Community version [here](https://docs.microsoft.com/en-us/visualstudio/releasenotes/vs2017-relnotes)

It is **crucial** to select [.NET desktop development] when being asked what packages to additionally install.

* DOTNET (download “.NET Core SDK”): “C:\Program Files\dotnet\”

DOWNLOAD [here](https://www.microsoft.com/net/download/windows)

* SAMTOOLS: Main directory "samtools" lies in "C:\Program Files".

DOWNLOAD [here](http://bow.codeplex.com/releases/view/87911)

Unzip the samtools.zip to "C:\Program Files".

* R version 3.1.2: Main directory "R" lies in "C:\Program Files".

DOWNLOAD [here](http://cran.cnr.berkeley.edu/bin/windows/base/old/3.1.2/)

* Rtools32

DOWNLOAD [here](https://cran.r-project.org/bin/windows/Rtools/)

## Supplemental Software

* Illumina® Variant Studio v3:

DOWNLOAD [here](https://illumina.account.box.com/login?redirect_url=https%3A%2F%2Fillumina.app.box.com%2Fv%2FVariantStudioV3)

* Integrative Genomics Viewer (IGV):

DOWNLOAD [here](https://software.broadinstitute.org/software/igv/download)

## Installation

1. Download and accept all default install locations.
2. Copy all relevant files and programs to their respective locations:

Go to “This PC” -> right-click This PC on left toolbar -> "Properties" -> "Advanced system settings" -> "Environment Variables..."

Add the following environment variables to the "Path" variable in the "System variables" window:

* C:\Rtools\bin
* C:\Rtools\gcc-4.6.3\bin
* C:\Program Files\R\R-3.1.2\bin
* C:\Users\***USER***\AppData\Local\Programs\Python\Python3#
* C:\Users\***USER*** \AppData\Local\Programs\Python\Python3#\Scripts
* C:\Program Files\Java\jre1.8.0\_161\bin
* C:\Program Files\Java\jdk1.8.0\_191\bin
* C:\Program Files\dotnet\
* C:\Program Files\samtools
* Rtools should be at top of list.
* Do not upgrade R and Rtools to a newer version
* If a newer version for JRE, JDK or Python was installed, update path accordingly.

1. Configure DECoN on the current machine:

Go to “C:\Gamidor\Appendix\DECoN-master\” and execute “setup.bat”

1. Install all relevant python packages from the requirements text file:

*pip install -U -r C:\Gamidor\MyScreen\requirements.txt*

1. MyScreen.exe can be moved to desktop directory if wanted

**Illumina® VariantStudio v3** [**Installation Guide**](https://www.illumina.com/content/dam/illumina-support/documents/documentation/software_documentation/variantstudio/Variant%20Studio_v3_Installation.pdf)

**IGV Installation Guide**

To start IGV from using Java Web Start:

1. Go to the IGV downloads page: http://www.broadinstitute.org/igv/download.
2. When prompted, register or log in as requested. You must register to download IGV.
3. Click the launch icon. The browser displays the web start launch window.
4. Select Open with Java™ Web Start and click OK. If the system displays messages about trusting the application, confirm that you trust the application. Web Start downloads and starts IGV.

Appendix B: Appendix Directories

## Version Analysis Files

* Main directory containing permanent genotyping and CNV installations and dependencies files lies in "C:\Gamidor\Appendix".
* An additional directory containing the following version dependent files lies in “C:\Gamidor\MyScreen\_Analysis\_v2.#\Script\docs”:
* Norm.RC.sort.txt
* Norm.RC.sort.vcf
* GenoAnno.pkl
* GenoBED.pkl
* ver2.agids
* CNVcustom.txt
* CNVtargets.bed
* CNVAnno.pkl
* CNVtargets.pkl
* Bedouin.AGID.pbz2
* word\_templates – directory with words hospital templates

“C:\Gamidor\MyScreen\_Analysis\_v2.#\Script\libs”:

* Python scripts and modules for current version

“C:\Gamidor\MyScreen\_Analysis\_v2.#\Script\images”:

* Used images

**\*\*\*\* If encountering error of “DLL not found” – you need to verify Visual Studio 2017 installation (**[**latest version of Visual Studio**](https://docs.microsoft.com/en-us/visualstudio/releasenotes/vs2017-relnotes)**). It is crucial to select [.NET desktop development] when being asked what packages to additionally install.**

Appendix C: Version Update Protocol

In order to correctly install a new version of MyScreen, the following steps must be taken:

1. Unzip the new version next to the current running version (in C:\Gamidor).
2. Move the entire current version folder into a folder named ‘Old’.

Appendix D: Legal Notices

## Methods and Limitations

**Targeted genotyping**

Targeted DNA mutation analysis is used to simultaneously determine the genotype of hundreds variants associated with many diseases, these variants (mutations) can increase the likelihood of conceiving a pregnancy with a hereditary condition. Most of the conditions on the panel are inherited in an autosomal recessive manner, meaning that both parents have to carry a mutation in the same disease gene in order to be at risk of having an affected child. Due to varying modes of inheritance and disease severity, there are a few diseases on the panel that can be transmitted when only one parent is a carrier.

**Sequencing**

The SureSelectXT HS Reagent Kits and protocol are used to prepare indexed library samples with molecular barcodes prior to target enrichment to allow high- sensitivity next- generation sequencing (NGS) on the Illumina platform.These regions are sequenced to high coverage and the sequences are compared to standards and references of normal variation. Mutations may not be detected in areas of lower coverage.

**Genotyping analysis and interpretation**

Only the mutations regions are compared and are genotyped (using Illumina® Pisces software).

The NGS technology test is highly reliable, with a >99% accuracy rate reported in the literature for targeted mutations. As with all medical screening tests, there is a chance of a false positive or false negative result. A “false positive” refers to the identification of a gene mutation that is not present. A “false negative” is the failure to recognize a mutation that indeed exists. Screening for the diseases on MyScreen panel may significantly reduce the likelihood of being a carrier but does not exclude the possibility of carrying another mutation within the genes of interest. The patient may still have a pathogenic variant that was not identified by this testing. Small insertions and deletions may not be as accurately determined as single nucleotide variants. Genes that have closely related pseudogenes are not well analyzed by this method.

**Copy number analysis**

Targeted large deletion and duplications analysis for several specific mutations is done by the DECoN software. Triplet repeats may not be detected.

*Fowler A, Mahamdallie S, Ruark E et al****. Accurate clinical detection of exon copy number variants in a targeted NGS panel using DECoN*** *[version 1; referees: 2 approved] Wellcome Open Research 2016, 1:20 (doi: 10.12688/wellcomeopenres.10069.1)*

SAM Tools provide various utilities for manipulating alignments in the SAM format, including sorting, merging, indexing and generating alignments in a per-position format.

*Li H.\*, Handsaker B.\*, Wysoker A., Fennell T., Ruan J., Homer N., Marth G., Abecasis G., Durbin R. and 1000 Genome Project Data Processing Subgroup (2009) The Sequence alignment/map (SAM) format and SAMtools. Bioinformatics, 25, 2078-9. [PMID: 19505943]*

**Testing limitations**

This test was developed and its performance characteristics determined by Gamidor Diagnostics Ltd (www.gamidor.com). It has not been cleared or approved by the FDA. Noteworthy, to date, the FDA does not require this test to go through premarket FDA review. Although molecular tests are highly accurate, it is a screening test only and not a diagnostic one. Rare analytic errors may occur that interfere with reporting. Sources of these errors include sample mix-up, trace contamination, bone marrow transplantation, blood transfusions and technical errors. The presence of additional variants nearby may interfere with mutation detection. The MyScreen test results must always be interpreted by a medical geneticist, genetic counselor or other qualified clinician in the context of clinical, familial and ancestral data. Genetic counseling is recommended to properly review and explain these results to the tested individual.

**Research**

Gamidor Diagnostics Ltd may use the aggregated and anonymous information provided by our users for internal research purposes.

**Legal Agreement**

You may not permit anyone else under your reasonable authority to copy, modify, create a derivative work of, reverse engineer, decompile or otherwise attempt to extract the source code or other basis of technology.

Any legal action or proceeding between Gamidor Diagnostics Ltd and you related to this agreement will be governed exclusively by the laws of the state of Israel. Any dispute, claim, or controversy in connection with or arising under the use of our Service or this agreement, its construction, existence, interpretation, validity, or any breach hereof which cannot be amicably settled between the parties, shall be finally and exclusively resolved by binding arbitration under the Rules of Arbitration of the Israeli Arbitration Association then prevailing.

**Limited Product Warranty**

Gamidor Diagnostics Ltd. warrants that this product will meet the specifications stated above. If any component of this product does not conform to these specifications Gamidor Diagnostics s Ltd. will at its sole discretion, as its sole and exclusive liability and as the users’ sole and exclusive remedy, replace the product at no charge or refund the cost of the product; provided that notice of non-conformance is given to Gamidor Diagnostics Ltd.within sixty (30) days of receipt of the product.

This warranty limits Gamidor Diagnostics Ltd’s liability to the replacement of this product or refund of the cost of the product. NO OTHER WARRANTIES OF ANY KIND, EXPRESS OR IMPLIED, INCLUDING WITHOUT LIMITATION IMPLIED WARRANTY OF MERCHANTABILITY OR FITNESS FOR A PARTICULAR PURPOSE OR NON-INFRINGEMENT, ARE PROVIDED BY GAMIDOR DIAGNOSTICS LTD. Gamidor Diagnostics Ltd. shall have no liability for any direct, indirect, consequential or incidental damages arising out of the use, the results of use or the inability to use this product and its components.

In no event shall Gamidor Diagnostics Ltd. be liable for claims for any other damages, whether direct, incidental, foreseeable, consequential, or special (including but not limited to loss of use, revenue or profit), whether based upon warranty, contract, tort (including negligence) or strict liability arising in connection with the sale or use or the failure of Gamidor Diagnostics Ltd. products to perform in accordance with the stated specifications.

Some components of nucleic acid analysis, such as specific methods and compositions for manipulating or visualizing nucleic acids for analysis, may be covered by one or more patents owned by other parties. Similarly, nucleic acids containing specific nucleotides sequences may be patented. Making, using, offering for sale, or selling such components or nucleic acids may require one or more licenses. Nothing in this document should be construed as an authorization or implicit license to make, use or sell any so covered component or nucleic acid under any such patents.